## Sheet1
ID | Description | Date of test | Test site | Picture | Material | Process | Creator (contact info and company) | Link to stl and cad files | Test 1: Expert evaluation | Test 2: Collection sufficiency | Test 3: PCR compatibility | Notes
This is a unique identifier, ideally including the source's name |  |  | E.g., BIDMC | Please take a high-resolution picture of the entire item provided. | Please describe all (known) materials in the bulb head, including plastic, any adhesive, and any flock | Please include a brief description of the design process that an expert would understand well enough to reproduce. For repurposed swabs, write "commercial." | How to contact the creator for questions/further development | Please send .stl and CAD files, which we will post. | If passes, write "Pass". If fails, write "Fail" and give a brief reason why. Should be filled out by someone who has experience actually using NP swabs in a clinical context. Please leave the name of this person. | If a Gram stain of the (inner) cheek made from the swab shows material broadly consistent with our control, then write "Pass". If not, "Fail" and give reason (too abrasive, did not pick up material, etc.) Note: BIDMC protocol does not include adding fluid. | As above, "Pass" vs. "Fail". Please indicate the protocol used. You can post the protocol in the Testing/ directory, and link to it here
Puritan Sterile polyester-tipped applicators 25-806 1PD |  | 2020-03-22 | BIDMC | img/Puritan/puritan_sterile_polyester_tipped_applicator_25-805_1PD.png | Polyester tip | Commercial | Puritan | NA | Did not pass (tip diameter too large) | Not done
Fisher polyester-tipped applicators 22363164 |  | 2020-03-22 | BIDMC | img/Fisher/fisher_polyester_tipped_applicators_22363164.png | Polyester tip | Commercial | Fisher | NA | Did not pass (tip diameter too large) | Not done
Puritan sterile foam tipped applicator 25-1506 1PF |  | 2020-03-22 | BIDMC | img/Puritan/puritan_sterile_foam_tipped_applicator_25-1506_1PF.png |  | Commercial | Fisher | NA | Pass (but tip diameter large, would be difficult to use) | Pass | Pass
Copan BD ESwab collection and transport system 220532 |  | 2020-03-22 | BIDMC | img/Copan_BD/copan_BD_ESwab_collection_and_transport_system_220532.png |  | Commercial | Fisher | NA | Pass | Pass
Hologic Aptima LBL-06548 |  | 2020-03-22 | BIDMC | img/Hologic/hologic_aptima_LBL-06548.png img/Hologic/hologic_aptima_LBL-06850.png |  | Commercial | Fisher | NA |  | Pass | Pass | There are two swabs in this kit. Both pass
Copan ESwab 220245 |  | 2020-03-22 | BIDMC | NA |  | Commercial | Fisher | NA |  |  | Pass
Hologic Aptima LBL-06850 |  | 2020-03-22 | BIDMC | Pending |  | Commercial | Fisher | NA | Pass (but tip diameter large)
Jifei "Carbon" (internal ID, *not* Carbon3D the company) |  | 2020-03-22 | BIDMC | (meth)acrylated urethane photopolymer |  |  | OPT Industries (Jifei Ou) |  |  | Pass | Pass
Jifei Material 2 |  | 2020-03-22 | BIDMC | (meth)acrylated urethane photopolymer | Material 2 |  | OPT Industries (Jifei Ou) |  | NOT SWAB MATERIAL ONLY | NOT SWAB MATERIAL ONLY | Pass
Nilson |  | Mach 20,2020 | BIDMC |  | Rose-print1 | Printed usign E Guide 2.0, cleaned in 99% IPA, 2000 Flash cure with ottoflash unit |  |  |  |  | Pass
Formlabs - Honey Tip |  | 2020-03-22 | BIDMC | img/FormLabs/honeydipper.png |  |  | FormLabs (Craig Brody) |  | Did not pass (shaft needs to be smooth; material too brittle) | Pass | Pass
Formlabs - Spiral Tip |  | 2020-03-22 | BIDMC | Formlabs Surgical Guide img/FormLabs/surgical_guide_technical_data_sheet_en.pdf |  |  | FormLabs (Craig Brody) |  | Did not pass (shaft needs to be smooth; material too brittle) | Did not pass (material did not transfer to slide?) | Pass
Formlabs - Bristle Tip |  | 2020-03-22 | BIDMC | img/FormLabs/bristles.png |  |  | FormLabs (Craig Brody) |  | Did not pass (shaft needs to be smooth; material too brittle) | Did not pass (material did not transfer to slide?) | Pass
Formlabs - Dental LT clear (Bristle, Honey tips) |  | 2020-03-22 | BIDMC | img/FormLabs/DentalLT.jpg |  |  | FormLabs (Craig Brody) |  | Did not pass (tip needs to be more flexible - will abutt posterior nasopharynx, but flexibility of NP neck, handle/body still needs to be more smooth)
Formlabs - Dental SG (Fern tip) |  | 2020-03-22 | BIDMC | img/FormLabs/SG.jpg |  |  | FormLabs (Craig Brody) |  | Did not pass (tip needs to be more flexible - will abutt posterior nasopharynx, but flexibility of NP neck, handle/body still needs to be more smooth) |  | Pass
Formlabs - Grey Pro (Spiral, Honey, Bristle) |  | 2020-03-22 | BIDMC | img/FormLabs/GreyPro.jpg |  |  | FormLabs (Craig Brody) |  | Did not pass (material brittle, tip needs to be more flexible - will abutt posterior nasopharynx, but flexibility of NP neck, handle/body still needs to be more smooth)
Formlabs - Durable (Fern, Honey, Bristle) |  | 2020-03-22 | BIDMC | img/FormLabs/Durable.jpg |  |  | FormLabs (Craig Brody) |  | Did Not Pass (Flexibility of tip appropriate, but neck for these designs is too elastic, difficult to pass middle turbinate, and folds in external nasal cavity) |  | Pass (All but Bristle tip -very few cells with that tip)
Formlabs Tough (Spiral, Bristle, Honey, Fern) |  | 2020-03-23 | BIDMC | img/FormLabs/tough.jpg |  |  | FormLabs (Craig Brody) |  | Did not pass (tip needs to be more flexible - will abutt posterior nasopharynx)
Formlabs Dental Model (Honey, Bristle) |  | 2020-03-24 | BIDMC | img/FormLabs/model.jpg |  |  | FormLabs (Craig Brody) |  | Did not pass (tip needs to be more flexible - will abutt posterior nasopharynx, and material too brittle)
Formlabs Tough 1500 (Spiral, Brittle) |  | 2020-03-25 | BIDMC | img/FormLabs/tough1500.jpg |  |  | FormLabs (Craig Brody) |  | Did not pass (tip needs to be more flexible - will abutt posterior nasopharynx)
Formlabs Tough - Tip 4 (Spiral, Bristle, Honey, Fern) |  | 2020-03-26 | BIDMC | img/FormLabs/tough.jpg |  |  | FormLabs (Craig Brody) |  | Did not pass (tip needs to be more flexible - will abutt posterior nasopharynx)
Harvard Army |  | 2020-03-22 | BIDMC | img/Harvard_army/new_skin.png |  |  | Army (Kit Parker) |  | Did not pass (material too brittle) | "New Skin" passes; others did not pass (too little material) |  | (3 glues used - Locktite, Dermabond, New skin)
Harvard Army Ultem |  | 2020-03-23 |  | img/Harvard_army/ultem.jpg |  |  | Army (Kit Parker) |  | Did not pass (material too brittle, tip inflexible
Harvard Army Onyx |  | 2020-03-23 |  | img/Harvard_army/onyx.jpg |  |  | Army (Kit Parker) |  | Did not pass (tip too inflexible)
Harvard Army Polyjet |  | 2020-03-23 |  |  |  |  | Army (Kit Parker) |  | Pass
W1 |  | 2020-03-22 | BIDMC | img/Wyss/wyss_w1.png | rayon, fiber, silicon glue, polycarbonate probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
W2 |  | 2020-03-22 | BIDMC | img/Wyss/wyss_w2.png | nylon fiber (0.75µm), silicon glue, polycarbonate probe |  | Wyss Institute (Adama Sesay) |  | Did not pass (needs a tapered, flexible end) | Did not pass (No visible cells or bacteria)
W3 |  | 2020-03-22 | BIDMC | img/Wyss/wyss_w3.png | rayon fiber, silicon glue, nylon probe |  | Wyss Institute (Adama Sesay) |  | Pass | Did not pass (very few cells/bacteria)
W4 |  | 2020-03-22 | BIDMC | img/Wyss/wyss_w4.png | rayon immfiber, silicon glue, commercial straw |  | Wyss Institute (Adama Sesay) |  | Did not pass (needs a tapered, flexible end) | Pass
W5 |  | 2020-03-22 | BIDMC | img/Wyss/wyss_w5.png | nylon fiber (0.75µm), nylon probe, silicon glue |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
W6 |  | 2020-03-22 | BIDMC | img/Wyss/wyss_w6.png | polycarbonate 3D-printed probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
W7 |  | 2020-03-22 | BIDMC | img/Wyss/wyss_w7.png | nylon 3D-printed probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
W8 |  | 2020-03-22 | BIDMC | img/Wyss/wyss_w8.png | nylon fiber (0.75µm), silicon glue, polycarbonate probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
VWR 1ul loop + pellon p45 fabric knot - Jenny |  | 2020-03-22 | HH (Seattle) | https://github.com/HMS-RIC/Covid19-NP-Swab/tree/master/img/Jenny/DIYswab | polystyrene loop and pellon P45 fabric | see https://github.com/HMS-RIC/Covid19-NP-Swab/wiki/DIY-swab-assembly | Jenny |  | Pass (doctor/swabber was able to use 5mm loop, newer iteration is <3mm) - Jenny | need review of smear pics
USF_1_1 | Rings w/ medium spacing distance | 2020-03-23 | BIDMC | img/USF/USF1-1.jpg |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass (tip is too sharp but appropriately flexible, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated | Pass
USF_1_2 | Rings w/ low spacing distance |  |  | img/USF/USF1-2.jpg |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass (tip is too sharp but stiff/brittle, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated | Pass
USF_1_3 | Rounded tip w/ protrusions |  |  | img/USF/USF1-3.jpg |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass (tip is too sharp and stiff, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated | Pass
USF_1_4 | Screw; finely threaded |  |  | img/USF/USF1-4.jpg |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass (tip is too sharp and stiff/brittle, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated | Pass
USF_1_5 | Screw; medium thread |  |  | img/USF/USF1-5.jpg |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass (tip is too sharp but we like the spring and how flexible it is, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated | Pass
USF_1_6 | Screw;large thread | 2020-03-23 |  | img/USF/USF1-6.jpg |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass (tip is too sharp but we like the spring and how flexible it is, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated | Pass
HP_H | HP swabs with holes | March 23,2020 | BIDMC | img/HP/HPH.jpg |  |  | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | **Conflicting Results** Pass per Jenny Hu 032320: doctor swabber would like the shaft to be longer but likes the shaft flexibility and tip size. Not pass per BIDMC: tip not very flexible and too sharp at the end | **Conflicting Results** Pass per Jenny Hu 032320 (collected a few cells/bacteria from cheek swab then rubbing directly on slides; better results from rinsing the tip with 150ul of buffer then transfering liquid to slide). Not pass per BIDMC protocol (no elution); very few cells collected |  | Doing some more mechanical measurements/testing this afternoon then my doctor friend will stop by after work and I'll do NP swabs with Copan, HP_H, HP_NH, VWRloop+facric. Will smear swabs directly on slides then try rinsing and looking at rinse liquid for cells/bacteria.
HP_NH | HP swabs without holes | 2020-03-23 | BIDMC | img/HP/HPNP.jpg |  |  | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | **Conflicting Results** Pass per Jenny Hu 032320: doctor swabber would like the shaft to be longer but likes the shaft flexibility and tip size. Not pass per BIDMC: tip not very flexible and too sharp at the end | **Conflicting Results** Pass per Jenny Hu 032320 (collected a few cells/bacteria from cheek swab then rubbing directly on slides; better results from rinsing the tip with 150ul of buffer then transfering liquid to slide). Not pass per BIDMC protocol (no elution); very few cells collected |  | Sending HP_NH to UW virology for PCR this afternoon I think a rinse/elution step is needed for the sample collection test since many of these designs shouldn't be expected to transfer well to a slide but can still elute into the VTM-Jenny
C_1 | Carbon 3D bag 1 | 2020-03-23 | BIDMC | img/Carbon/C1.jpg |  |  | Carbon 3D ( Hardik Kabaria <hardik@carbon3d.com) |  | Pass (good stiffness of handle, and flexible tip), tapered more flexible neck would be better | Pass
C_2 | Carbon 3D bag 2 | 2020-03-23 | BIDMC | img/Carbon/C2.jpg |  |  | Carbon 3D (Ramy Arnaout; rarnaout@bidmc.harvard.edu) |  | Pass (good stiffness of material, and flexible tip), tapered more flexible neck would be better | Pass
P_0 | Swab_test_Wyss | 2020-03-23 | BIDMC | img/Gorelik/P0.jpg | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Did not pass - tip needs to be more flexible, but stiffness of neck and elasticity of material appropriate)
P_1 | Solid_Circular | 2020-03-23 | BIDMC | img/Gorelik/P1.jpg | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Pass (tip more appropriately flexible, but would make tip end less jagged)
P_2 | Spring_Circular | 2020-03-23 | BIDMC | img/Gorelik/P2.jpg | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Pass (tip more appropriately flexible, but would make tip end less jagged) | Pass
P_3 | Spring_Flat | 2020-03-23 | BIDMC | img/Gorelik/P3.jpg | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Did not pass - tip needs to be more flexible, but stiffness of neck and elasticity of material appropriate) | Pass
P_4 | Solid_Flat | 2020-03-23 | BIDMC | img/Gorelik/P4.jpg | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Pass (tip more appropriately flexible, but would make tip end less jagged)
P_5 | Swab_Concentric | 2020-03-23 | BIDMC | img/Gorelik/P5.jpg | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Pass (tip more appropriately flexible, but would make tip end less jagged) | Pass
P_6 | Swab_concentric_bigger_pitch | 2020-03-23 | BIDMC | img/Gorelik/P6.jpg | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Pass (tip more appropriately flexible, but would make tip end less jagged) | Pass
P_7 | Swab_test_head_4_wyss | 2020-03-23 | BIDMC | img/Gorelik/P7.jpg | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Pass for tip head
J_1 | flexible swab w/ case | March 24th, 2020 | BIDMC | img/OPT_industries/J_1.jpg | methylacrylic urethane |  | opticalindustries, jifei jifei@optindustries.com |  | Pass but design too short | Pass
USF_2_1 | swab w/ fine nubs | March 24th, 2020 | BIDMC | img/USF/USF_2_1.jpg |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass (tip too sharp and should be more flexible, but body design is appropriate) | Pass
USF_2_2 | swab w/ finely spaced rings | Marh 24th, 2020 | BIDMC | img/USF/USF_2_2.jpg |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass (tip too sharp and but apporpirately flexible, but body design is appropriate) | Pass
B_1 | swab w/very fine nubs | March 24th, 2020 | BIDMC | img/Baker/B_1 |  |  | Daniel Baker |  | Did not pass (tip is too sharp and ideally more flexible tip, but body design is appropriate) | Pass
S_1 | Swab w/ thick thread | March 24th, 2020 | BIDMC | img/Seriano/S_1 |  |  | David Seriano |  | Did not pass (tip is too sharp, but apporpiately flexible and body design is appopriate) | Pass
S_2 | swab w/ fine spring | March 24th | BIDMC | img/Seriano/S_2 |  |  | David Seriano |  | Did not pass (tip is too sharp, but very flexible, perhaps too flexible, S1 may be better, and body design is appopriate) | Pass
S_3 | swab w/ fine thread | March 24th | BIDMC | img/Seriano/S_3 |  |  | David Seriano |  | Did not pass (tip is too sharp, and body design is appopriate) | Pass
S_4 | swab w/finely speaced rings | March 24th | BIDMC | img/Seriano/S_4 |  |  | David Seriano |  | Did not pass (tip is too sharp, and body design is appopriate) | Pass
S_5 | swab w/ low ring spacing | March 24th | BIDMC | img/Seriano/S_5 |  |  | David Seriano |  | Did not pass (tip is too sharp, and body design is appopriate) | Pass
S_6 | swab w/fine rings | March 24th | BIDMC | img/Seriano/S_6 |  |  | David Seriano |  |  | Pass
FT_1 | Foam Swab. Pt: HT1171B | March 24th, 2020 | BIDMC | Img/Foamtec/FT_1 |  |  | Foamtec International; Michael Strauss |  | Did not pass (diameter of tip too large, flexible bendable neck would be better) | Did not pass (not many cells collected)
FT_2 | Foam Swab. Pt: HT1501T2-5 | March 24th, 2020 | BIDMC | Img/Foamtec/FT_2 |  |  | Foamtec International; Michael Strauss |  | Did not pass (diameter of tip too large, flexible bendable neck would be better) | Did not pass (not many cells collected)
FT_3 | Foam. Pt. HT1501TC-5 | March 24th, 2020 | BIDMC | Img/Foamtec/FT_3 |  |  | Foamtec International; Michael Strauss |  | Did not pass (diameter of tip too large, flexible bendable neck would be better) | Did not pass (not many cells collected)
FT_4 | Foam Swab. Pt: HT1001-5 | March 24th, 2020 | BIDMC | Img/Foamtec/FT_4 |  |  | Foamtec International; Michael Strauss |  | Did not pass (diameter tip, diameter of tip too large, flexible bendable neck would be better) | Did not pass (not many cells collected)
EV_1 | DOWEL | March 24th, 2020 | BIDMC |  | PCA 2000. 10 minutes. 100% |  | Enviriontec. Al Siblani
EV_2 | DOWEL | March 24th, 2020 | BIDMC |  | PCA100 w/nitrogen | 2x10seconds w/nitrogen | Enviriontec. Al Siblani
EV_3 | DOWEL | March 24th, 2020 | BIDMC |  | PCA 100 w/nitrogen | 2x30seconds UV w/nitrogen | Enviriontec. Al Siblani
EV_4 | DOWEL | March 24th, 2020 | BIDMC |  | PCA 100 w/nitrogen | 2x20seconds w/nitrogen | Enviriontec. Al Siblani