## Sheet1
| Date of test | Test site | Picture | Material | Process | Creator (contact info and company) | Link to stl and cad files | Test 1: Expert evaluation | Test 2: Collection sufficiency | Test 3: PCR compatibility | Notes
This is a unique identifier, ideally including the source's name |  | E.g., BIDMC | Please take a high-resolution picture of the entire item provided. | Please describe all (known) materials in the bulb head, including plastic, any adhesive, and any flock | Please include a brief description of the design process that an expert would understand well enough to reproduce. For repurposed swabs, write "commercial." | How to contact the creator for questions/further development | Please send .stl and CAD files, which we will post. | If passes, write "Pass". If fails, write "Fail" and give a brief reason why. Should be filled out by someone who has experience actually using NP swabs in a clinical context. Please leave the name of this person. | If a Gram stain of the (inner) cheek made from the swab shows material broadly consistent with our control, then write "Pass". If not, "Fail" and give reason (too abrasive, did not pick up material, etc.) | As above, "Pass" vs. "Fail". Please indicate the protocol used. You can post the protocol in the Testing/ directory, and link to it here
Puritan Sterile polyester-tipped applicators 25-806 1PD | 2020-03-22 | BIDMC | img/Puritan/puritan_sterile_polyester_tipped_applicator_25-805_1PD.png | Polyester tip | Commercial | Puritan | NA | Did not pass (tip diameter too large) | Not done
Fisher polyester-tipped applicators 22363164 | 2020-03-22 | BIDMC | img/Fisher/fisher_polyester_tipped_applicators_22363164.png | Polyester tip | Commercial | Fisher | NA | Did not pass (tip diameter too large) | Not done
Puritan sterile foam tipped applicator 25-1506 1PF | 2020-03-22 | BIDMC | img/Puritan/puritan_sterile_foam_tipped_applicator_25-1506_1PF.png |  | Commercial | Fisher | NA | Pass (but tip diameter large, would be difficult to use) | Pass | Pass
Copan BD ESwab collection and transport system 220532 | 2020-03-22 | BIDMC | img/Copan_BD/copan_BD_ESwab_collection_and_transport_system_220532.png |  | Commercial | Fisher | NA | Pass | Pass
Hologic Aptima LBL-06548 | 2020-03-22 | BIDMC | img/Hologic/hologic_aptima_LBL-06548.png img/Hologic/hologic_aptima_LBL-06850.png |  | Commercial | Fisher | NA |  | Pass | Pass | There are two swabs in this kit. Both pass
Copan ESwab 220245 | 2020-03-22 | BIDMC | NA |  | Commercial | Fisher | NA |  |  | Pass
Hologic Aptima LBL-06850 | 2020-03-22 | BIDMC | Pending |  | Commercial | Fisher | NA | Pass (but tip diameter large)
Jifei Carbon | 2020-03-22 | BIDMC | img/OPT_industries/reinforced_cilia.png |  |  | OPT Industries (Jifei Ou) |  |  | Pass | Pass
Jifei Material 2 | 2020-03-22 | BIDMC |  | Material 2 |  | OPT Industries (Jifei Ou) |  | NOT SWAB MATERIAL ONLY | NOT SWAB MATERIAL ONLY | Pass
Formlabs - Honey | 2020-03-22 | BIDMC | img/FormLabs/honeydipper.png |  |  | FormLabs (Craig Brody) |  | Did not pass (shaft needs to be smooth; material too brittle) | Pass | Pass
Formlabs - Spiral | 2020-03-22 | BIDMC |  |  |  | FormLabs (Craig Brody) |  | Did not pass (shaft needs to be smooth; material too brittle) | Did not pass (material did not transfer to slide?) | Pass
Formlabs - Bristle | 2020-03-22 | BIDMC | img/FormLabs/bristles.png |  |  | FormLabs (Craig Brody) |  | Did not pass (shaft needs to be smooth; material too brittle) | Did not pass (material did not transfer to slide?) | Pass
Harvard Army | 2020-03-22 | BIDMC | img/Harvard_army/new_skin.png |  |  | Army (Kit Parker) |  | Did not pass (material too brittle) | "New Skin" passes; others did not pass (too little material) |  | (3 glues used - Locktite, Dermabond, New skin)
W1 | 2020-03-22 | BIDMC | img/Wyss/wyss_w1.png | rayon, fiber, silicon glue, polycarbonate probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
W2 | 2020-03-22 | BIDMC | img/Wyss/wyss_w2.png | nylon fiber (0.75µm), silicon glue, polycarbonate probe |  | Wyss Institute (Adama Sesay) |  | Did not pass (needs a tapered, flexible end) | Did not pass (No visible cells or bacteria)
W3 | 2020-03-22 | BIDMC | img/Wyss/wyss_w3.png | rayon fiber, silicon glue, nylon probe |  | Wyss Institute (Adama Sesay) |  | Pass | Did not pass (very few cells/bacteria)
W4 | 2020-03-22 | BIDMC | img/Wyss/wyss_w4.png | rayon immfiber, silicon glue, commercial straw |  | Wyss Institute (Adama Sesay) |  | Did not pass (needs a tapered, flexible end) | Pass
W5 | 2020-03-22 | BIDMC | img/Wyss/wyss_w5.png | nylon fiber (0.75µm), nylon probe, silicon glue |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
W6 | 2020-03-22 | BIDMC | img/Wyss/wyss_w6.png | polycarbonate 3D-printed probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
W7 | 2020-03-22 | BIDMC | img/Wyss/wyss_w7.png | nylon 3D-printed probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
W8 | 2020-03-22 | BIDMC | img/Wyss/wyss_w8.png | nylon fiber (0.75µm), silicon glue, polycarbonate probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
VWR 1ul loop + pellon p45 fabric knot - Jenny | 2020-03-22 | HH (Seattle) | https://github.com/HMS-RIC/Covid19-NP-Swab/tree/master/img/Jenny/DIYswab | polystyrene loop and pellon P45 fabric | see https://github.com/HMS-RIC/Covid19-NP-Swab/wiki/DIY-swab-assembly | Jenny |  | Pass (doctor/swabber was able to use 5mm loop, newer iteration is <3mm) | need review of smear pics